## PDR plus
| TG-43 CONSENSUS |  |  |  |  | GammaMed PDR 192Ir Plus
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.123 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL( r) |  |  |  | F(r,q) | [18 empty] | QA Along-away
| L = | 0.05 | cm |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL(r) L = 0,05 cm unbounded |  | q (º) | 0 | 0.2 | 0.4 | 0.6 | 0.8 | 1 | 1.25 | 1.5 | 1.75 | 2 | 2.5 | 3 | 3.5 | 4 | 5 | 6 | 8 | 10 | z (cm) | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 1.0046 |  | 0 | 0.9403 | 0.9403 | 0.94 | 0.9397 | 0.948 | 0.951 | 0.9523 | 0.952 | 0.952 | 0.952 | 0.952 | 0.952 | 0.952 | 0.954 | 0.9535 | 0.96 | 0.96 | 0.965 | 7 | 0.021987 | 0.022158 | 0.022097 | 0.021994 | 0.021865 | 0.021437 | 0.020823 | 0.019101 | 0.017037 | 0.014855 | 0.012812 | 0.010977
| 0.2 | 1.0046 |  | 1 | 0.941 | 0.941 | 0.941 | 0.941 | 0.949 | 0.953 | 0.953 | 0.9526 | 0.953 | 0.953 | 0.953 | 0.953 | 0.953 | 0.9544 | 0.956 | 0.9607 | 0.9614 | 0.9691 | 6 | 0.030174 | 0.030374 | 0.030253 | 0.030084 | 0.029817 | 0.029028 | 0.027891 | 0.024902 | 0.021529 | 0.018256 | 0.015322 | 0.012824
| 0.25 | 0.99992 |  | 2 | 0.9464 | 0.9464 | 0.9435 | 0.9406 | 0.949 | 0.953 | 0.953 | 0.953 | 0.953 | 0.954 | 0.955 | 0.957 | 0.958 | 0.9582 | 0.9588 | 0.9678 | 0.9697 | 0.9758 | 5 | 0.043366 | 0.043613 | 0.043534 | 0.043277 | 0.04276 | 0.041123 | 0.038831 | 0.033275 | 0.027542 | 0.02248 | 0.018279 | 0.014888
| 0.5 | 0.99431 |  | 3 | 0.945 | 0.945 | 0.945 | 0.945 | 0.9513 | 0.953 | 0.953 | 0.9534 | 0.954 | 0.955 | 0.96 | 0.9603 | 0.96 | 0.9601 | 0.9617 | 0.9685 | 0.9697 | 0.9755 | 4 | 0.067883 | 0.068103 | 0.067696 | 0.066883 | 0.065624 | 0.061843 | 0.056782 | 0.04553 | 0.035515 | 0.0276 | 0.021593 | 0.017097
| 0.75 | 0.99768 |  | 4 | 0.9481 | 0.9481 | 0.9486 | 0.9491 | 0.9549 | 0.957 | 0.957 | 0.9569 | 0.957 | 0.958 | 0.962 | 0.9619 | 0.961 | 0.9613 | 0.9639 | 0.9691 | 0.9706 | 0.976 | 3 | 0.12028 | 0.1208 | 0.11912 | 0.11641 | 0.11231 | 0.10085 | 0.087643 | 0.063479 | 0.045657 | 0.033422 | 0.025071 | 0.019263
| 1 | 1 |  | 5 | 0.9478 | 0.9478 | 0.9493 | 0.9508 | 0.9555 | 0.9571 | 0.9573 | 0.9574 | 0.957 | 0.959 | 0.962 | 0.963 | 0.961 | 0.9613 | 0.9651 | 0.9694 | 0.9714 | 0.9765 | 2 | 0.26941 | 0.26805 | 0.26046 | 0.24584 | 0.22575 | 0.18189 | 0.14258 | 0.087911 | 0.057231 | 0.039259 | 0.028283 | 0.021152
| 1.5 | 1.0032 |  | 6 | 0.9467 | 0.9467 | 0.949 | 0.9513 | 0.9556 | 0.9568 | 0.9581 | 0.9587 | 0.958 | 0.96 | 0.962 | 0.9632 | 0.963 | 0.9632 | 0.9675 | 0.9706 | 0.9725 | 0.9773 | 1.5 | 0.47689 | 0.47141 | 0.4449 | 0.39991 | 0.34797 | 0.25267 | 0.18238 | 0.10155 | 0.06266 | 0.041794 | 0.029599 | 0.021898
| 2 | 1.0076 |  | 7 | 0.9476 | 0.9476 | 0.9514 | 0.9552 | 0.9591 | 0.9597 | 0.9616 | 0.9623 | 0.962 | 0.9616 | 0.963 | 0.963 | 0.9627 | 0.9658 | 0.9702 | 0.9722 | 0.9738 | 0.978 | 1 | 1.0689 | 1.0316 | 0.89567 | 0.7211 | 0.56515 | 0.3493 | 0.22766 | 0.11412 | 0.067171 | 0.043799 | 0.03061 | 0.022457
| 3 | 1.0123 |  | 8 | 0.9507 | 0.9507 | 0.9547 | 0.9587 | 0.9618 | 0.9617 | 0.9639 | 0.9644 | 0.964 | 0.9638 | 0.964 | 0.965 | 0.9649 | 0.9676 | 0.9724 | 0.9739 | 0.9756 | 0.9791 | 0.5 | 4.2092 | 3.5611 | 2.2515 | 1.3888 | 0.90382 | 0.45316 | 0.2671 | 0.12319 | 0.070163 | 0.045072 | 0.031234 | 0.022797
| 4 | 1.0136 |  | 9 | 0.9505 | 0.9505 | 0.9551 | 0.9597 | 0.9626 | 0.9622 | 0.965 | 0.9655 | 0.965 | 0.9653 | 0.967 | 0.967 | 0.9667 | 0.9689 | 0.9738 | 0.9748 | 0.9767 | 0.9799 | 0 | 3819734318.8827 | 17.9107 | 4.4637 | 1.9915 | 1.123 | 0.50074 | 0.28293 | 0.12634 | 0.071153 | 0.045479 | 0.031431 | 0.022903
| 5 | 1.0123 |  | 10 | 0.956 | 0.956 | 0.9605 | 0.965 | 0.9672 | 0.9662 | 0.969 | 0.969 | 0.969 | 0.9688 | 0.97 | 0.9712 | 0.9697 | 0.9713 | 0.976 | 0.9765 | 0.9781 | 0.9805 | -0.5 | 4.2593 | 3.526 | 2.2383 | 1.3838 | 0.90204 | 0.45236 | 0.267 | 0.12309 | 0.070157 | 0.04505 | 0.031227 | 0.022792
| 6 | 1.0074 |  | 15 | 0.9686 | 0.9686 | 0.9728 | 0.977 | 0.9784 | 0.9768 | 0.9799 | 0.9786 | 0.9796 | 0.9796 | 0.9799 | 0.9804 | 0.981 | 0.981 | 0.9845 | 0.984 | 0.9862 | 0.9873 | -1 | 1.0493 | 1.0101 | 0.88589 | 0.71558 | 0.56263 | 0.34807 | 0.2272 | 0.11397 | 0.067144 | 0.043758 | 0.030595 | 0.022446
| 8 | 0.9878 |  | 20 | 0.9729 | 0.9729 | 0.9853 | 0.9884 | 0.9895 | 0.9885 | 0.9905 | 0.9889 | 0.9915 | 0.9899 | 0.9887 | 0.9898 | 0.9906 | 0.9898 | 0.9911 | 0.9902 | 0.9921 | 0.9921 | -1.5 | 0.45893 | 0.45296 | 0.43858 | 0.39531 | 0.34494 | 0.25124 | 0.18179 | 0.10133 | 0.062603 | 0.041742 | 0.029573 | 0.021882
| 10 | 0.9536 |  | 30 | 0.9794 | 0.9794 | 0.9953 | 0.9981 | 0.9989 | 0.9993 | 0.9992 | 0.9998 | 1.0002 | 0.9999 | 0.9986 | 1.0004 | 0.9996 | 0.9995 | 0.9993 | 0.9986 | 0.9993 | 0.9988 | -2 | 0.25433 | 0.25108 | 0.25524 | 0.24273 | 0.22343 | 0.18064 | 0.14186 | 0.087682 | 0.057132 | 0.039195 | 0.028247 | 0.02113
|  |  |  | 40 | 0.985 | 0.985 | 0.9995 | 1.0032 | 1.0035 | 1.0034 | 1.003 | 1.0035 | 1.0032 | 1.0036 | 1.0028 | 1.0037 | 1.0018 | 1.0022 | 1.0019 | 1.0017 | 1.0015 | 1.0019 | -3 | 0.10914 | 0.10771 | 0.11479 | 0.11433 | 0.11084 | 0.099761 | 0.086939 | 0.063198 | 0.045471 | 0.033332 | 0.025017 | 0.019229
|  |  |  | 50 | 0.99 | 0.99 | 1.0032 | 1.0054 | 1.006 | 1.0051 | 1.0044 | 1.004 | 1.0048 | 1.0044 | 1.004 | 1.0042 | 1.0028 | 1.004 | 1.0038 | 1.0036 | 1.0027 | 1.0029 | -4 | 0.058982 | 0.059239 | 0.06406 | 0.065126 | 0.064564 | 0.061092 | 0.056243 | 0.045243 | 0.035347 | 0.027497 | 0.021527 | 0.017055
|  |  |  | 60 | 0.9916 | 0.9916 | 1.0045 | 1.0055 | 1.0069 | 1.0062 | 1.0051 | 1.0059 | 1.0057 | 1.0048 | 1.0048 | 1.0044 | 1.0035 | 1.0049 | 1.0041 | 1.0033 | 1.0026 | 1.0032 | -5 | 0.036112 | 0.037606 | 0.040461 | 0.041451 | 0.041538 | 0.040484 | 0.038375 | 0.03303 | 0.02739 | 0.022374 | 0.018207 | 0.014839
|  |  |  | 70 | 0.9931 | 0.9931 | 1.004 | 1.0051 | 1.0057 | 1.0053 | 1.004 | 1.0056 | 1.0045 | 1.0035 | 1.0037 | 1.0035 | 1.003 | 1.0039 | 1.0033 | 1.0027 | 1.0024 | 1.0029 | -6 | 0.02386 | 0.025618 | 0.027693 | 0.028543 | 0.028878 | 0.028525 | 0.027527 | 0.024667 | 0.021375 | 0.018149 | 0.015244 | 0.01277
|  |  |  | 80 | 0.9992 | 0.9992 | 1.0018 | 1.0024 | 1.0027 | 1.0028 | 1.0025 | 1.0035 | 1.0028 | 1.0021 | 1.0021 | 1.0022 | 1.002 | 1.002 | 1.0019 | 1.0017 | 1.0016 | 1.0017 | -7 | 0.016586 | 0.018519 | 0.020068 | 0.020744 | 0.021052 | 0.020969 | 0.020505 | 0.018885 | 0.016892 | 0.014747 | 0.012729 | 0.010916
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 100 | 0.9975 | 0.9975 | 1.0024 | 1.003 | 1.0022 | 1.0023 | 1.0023 | 1.0022 | 1.0019 | 1.0019 | 1.0014 | 1.0013 | 1.0012 | 1.0019 | 1.001 | 1.0012 | 1.001 | 1.0007
|  |  |  | 110 | 0.9923 | 0.9923 | 1.0026 | 1.0053 | 1.0043 | 1.0042 | 1.0032 | 1.0037 | 1.0027 | 1.003 | 1.0025 | 1.0022 | 1.0014 | 1.0033 | 1.0018 | 1.0013 | 1.001 | 1.0013
|  |  |  | 120 | 0.989 | 0.989 | 1.0009 | 1.0032 | 1.0033 | 1.0036 | 1.0027 | 1.004 | 1.0026 | 1.0022 | 1.0021 | 1.0018 | 1.0011 | 1.003 | 1.0018 | 1.0007 | 1.0001 | 1.0007
|  |  |  | 130 | 0.9848 | 0.9848 | 0.9981 | 1.0014 | 1.0007 | 1.001 | 1.0009 | 1.0021 | 1.0005 | 0.9994 | 1.0005 | 0.9995 | 0.9998 | 1.0002 | 0.9989 | 0.9997 | 0.9987 | 0.9988
|  |  |  | 140 | 0.9773 | 0.9773 | 0.9937 | 0.9967 | 0.9959 | 0.9962 | 0.9963 | 0.9971 | 0.9972 | 0.9966 | 0.9969 | 0.9977 | 0.9963 | 0.9977 | 0.9961 | 0.9967 | 0.9956 | 0.9948
|  |  |  | 150 | 0.9696 | 0.9696 | 0.9874 | 0.99 | 0.9903 | 0.9899 | 0.9894 | 0.99 | 0.9897 | 0.9909 | 0.9896 | 0.9913 | 0.9896 | 0.9912 | 0.9919 | 0.991 | 0.9909 | 0.9901
|  |  |  | 160 | 0.956 | 0.956 | 0.9721 | 0.9753 | 0.975 | 0.9752 | 0.9771 | 0.977 | 0.977 | 0.977 | 0.9765 | 0.9777 | 0.9777 | 0.9778 | 0.9781 | 0.9783 | 0.9793 | 0.9807
|  |  |  | 165 | 0.9556 | 0.9556 | 0.9579 | 0.9602 | 0.9599 | 0.96 | 0.961 | 0.962 | 0.9624 | 0.9626 | 0.9653 | 0.9663 | 0.9672 | 0.9686 | 0.9672 | 0.9696 | 0.9698 | 0.9719
|  |  |  | 170 | 0.9375 | 0.9375 | 0.937 | 0.9365 | 0.9334 | 0.9326 | 0.934 | 0.9346 | 0.936 | 0.937 | 0.9386 | 0.939 | 0.939 | 0.9443 | 0.9437 | 0.9489 | 0.954 | 0.956
|  |  |  | 171 | 0.925 | 0.925 | 0.925 | 0.925 | 0.9247 | 0.9232 | 0.9206 | 0.9246 | 0.9239 | 0.9246 | 0.9281 | 0.9291 | 0.9284 | 0.935 | 0.9357 | 0.941 | 0.9486 | 0.951
|  |  |  | 172 | 0.908 | 0.908 | 0.91 | 0.912 | 0.912 | 0.9117 | 0.912 | 0.9121 | 0.912 | 0.9122 | 0.9157 | 0.9177 | 0.9166 | 0.9251 | 0.9275 | 0.9327 | 0.9421 | 0.9447
|  |  |  | 173 | 0.896 | 0.896 | 0.897 | 0.898 | 0.899 | 0.899 | 0.899 | 0.8991 | 0.899 | 0.8993 | 0.9026 | 0.903 | 0.903 | 0.9126 | 0.9163 | 0.9209 | 0.9325 | 0.9374
|  |  |  | 174 | 0.88 | 0.88 | 0.88 | 0.88 | 0.88 | 0.881 | 0.8832 | 0.8833 | 0.883 | 0.884 | 0.885 | 0.8869 | 0.887 | 0.8966 | 0.9031 | 0.9081 | 0.9217 | 0.9289
|  |  |  | 175 | 0.8227 | 0.8227 | 0.8318 | 0.8409 | 0.85 | 0.8591 | 0.8594 | 0.8585 | 0.859 | 0.8588 | 0.8632 | 0.8652 | 0.8659 | 0.8776 | 0.8872 | 0.892 | 0.9058 | 0.9171
|  |  |  | 176 | 0.8136 | 0.8136 | 0.8152 | 0.8168 | 0.8184 | 0.82 | 0.822 | 0.823 | 0.8242 | 0.8312 | 0.8362 | 0.8363 | 0.8398 | 0.8513 | 0.8656 | 0.8716 | 0.8865 | 0.9037
|  |  |  | 177 | 0.7436 | 0.7436 | 0.7492 | 0.7548 | 0.7604 | 0.766 | 0.773 | 0.78 | 0.7795 | 0.79 | 0.7972 | 0.798 | 0.8039 | 0.8143 | 0.8357 | 0.8444 | 0.86 | 0.8844
|  |  |  | 178 | 0.5676 | 0.5676 | 0.5812 | 0.5948 | 0.6084 | 0.622 | 0.639 | 0.656 | 0.673 | 0.69 | 0.724 | 0.729 | 0.739 | 0.7559 | 0.7861 | 0.7989 | 0.8163 | 0.8515
|  |  |  | 179 | 0.85498 | 0.85498 | 0.85236 | 0.84974 | 0.84712 | 0.8445 | 0.84123 | 0.83795 | 0.83468 | 0.8314 | 0.82485 | 0.8183 | 0.81175 | 0.8052 | 0.7921 | 0.779 | 0.7528 | 0.8019
|  |  |  | 180 | 0.96152 | 0.96152 | 0.95454 | 0.94756 | 0.94058 | 0.9336 | 0.92488 | 0.91615 | 0.90743 | 0.8987 | 0.88125 | 0.8638 | 0.84635 | 0.8289 | 0.794 | 0.7591 | 0.6893 | 0.694