## info
General Information
Source Information
Source name: | CSM40
Radionuclide: | Cs-137
Model Number:
Distributed by: | E&Z Bebig
Manufactured by:
Calculation details
Method: | Monte Carlo
Code: | GEANT4 and Penelope
Version: | 4.5.0 (Geant4) 2008 (Penelope)
Phantom material: | MC: Liquid water
Phantom size: | MC: Sphere 40 cm in radius
Cutoff energy: | 10 keV
Cross-section: | EPDL97
Electron tracking: | No
Publication | Monte Carlo dosimetric study of the medium dose rate CSM40 source
J. Vijande, D. Granero, J. Pérez-Calatayud and F. Ballester
Applied Radiation and Isotopes 82 (2013) 283-288
http://dx.doi.org/10.1016/j.apradiso.2013.09.003

## Source
Source design and materials | 137Cs - LDR Bebig CSM40

## GL(r,t)
Geometry Function |  | 137Cs - LDR Bebig CSM40
Active source length, L= | 1.65 | cm

## Dose Rate Constant
Dose Rate Constant |  |  |  | 137Cs - LDR Bebig CSM40
| L = | 0.92653 | ± | 0.0005 | cGy h-1 U-1

## gL(r)
Radial Dose Function Parameters |  |  | 137Cs - LDR Bebig CSM40
Leff = | 1.65 | cm
r (cm) | gL(r )
0.15 | 1.0861
0.2 | 1.0409
0.25 | 1.0256
0.3 | 1.0183
0.35 | 1.014
0.4 | 1.0111
0.45 | 1.0091
0.5 | 1.0078
0.55 | 1.0064
0.6 | 1.0053
0.65 | 1.0046
0.7 | 1.0039
0.75 | 1.0032
0.8 | 1.0026
0.85 | 1.0017
0.9 | 1.0012
0.95 | 1.0006
1 | 1
1.05 | 0.99964
1.1 | 0.9991
1.15 | 0.99833
1.2 | 0.99787
1.25 | 0.99743
1.3 | 0.99676
1.35 | 0.99641
1.4 | 0.99572
1.45 | 0.9951
1.5 | 0.99468
1.55 | 0.99411
1.6 | 0.99381
1.65 | 0.99333
1.7 | 0.99291
1.75 | 0.99246
1.8 | 0.99193
1.85 | 0.99144
1.9 | 0.99094
1.95 | 0.99046
2 | 0.98986
2.05 | 0.98938
2.1 | 0.98879
2.15 | 0.98822
2.2 | 0.98775
2.25 | 0.98721
2.3 | 0.98674
2.35 | 0.98617
2.4 | 0.9857
2.45 | 0.9851
2.5 | 0.98444
2.55 | 0.98395
2.6 | 0.98336
2.65 | 0.98294
2.7 | 0.9824
2.75 | 0.98175
2.8 | 0.98116
2.85 | 0.98074
2.9 | 0.98009
2.95 | 0.97953
3 | 0.97906
3.05 | 0.97844
3.1 | 0.97783
3.15 | 0.97724
3.2 | 0.97683
3.25 | 0.97633
3.3 | 0.97571
3.35 | 0.97498
3.4 | 0.97436
3.45 | 0.97392
3.5 | 0.97326
3.55 | 0.97266
3.6 | 0.9721
3.65 | 0.97151
3.7 | 0.97102
3.75 | 0.97038
3.8 | 0.96971
3.85 | 0.96913
3.9 | 0.96853
3.95 | 0.96782
4 | 0.96718
4.05 | 0.96658
4.1 | 0.96596
4.15 | 0.96525
4.2 | 0.96467
4.25 | 0.96409
4.3 | 0.96342
4.35 | 0.9627
4.4 | 0.96201
4.45 | 0.96144
4.5 | 0.96075
4.55 | 0.96018
4.6 | 0.95951
4.65 | 0.95879
4.7 | 0.95816
4.75 | 0.95747
4.8 | 0.95674
4.85 | 0.95607
4.9 | 0.95526
4.95 | 0.95456
5 | 0.95389
5.05 | 0.95322
5.1 | 0.95266
5.15 | 0.95208
5.2 | 0.95141
5.25 | 0.95072
5.3 | 0.94993
5.35 | 0.9492
5.4 | 0.9485
5.45 | 0.94772
5.5 | 0.94696
5.55 | 0.94631
5.6 | 0.94564
5.65 | 0.94493
5.7 | 0.94418
5.75 | 0.94351
5.8 | 0.94277
5.85 | 0.94201
5.9 | 0.94128
5.95 | 0.94051
6 | 0.93979
6.05 | 0.93899
6.1 | 0.93829
6.15 | 0.93761
6.2 | 0.93689
6.25 | 0.93611
6.3 | 0.93547
6.35 | 0.93465
6.4 | 0.93384
6.45 | 0.93312
6.5 | 0.93244
6.55 | 0.93161
6.6 | 0.9309
6.65 | 0.9301
6.7 | 0.92931
6.75 | 0.92861
6.8 | 0.92778
6.85 | 0.92698
6.9 | 0.92606
6.95 | 0.92532
7 | 0.92464
7.05 | 0.92377
7.1 | 0.92297
7.15 | 0.92222
7.2 | 0.92138
7.25 | 0.92058
7.3 | 0.91983
7.35 | 0.91901
7.4 | 0.91812
7.45 | 0.91728
7.5 | 0.91649
7.55 | 0.91563
7.6 | 0.91483
7.65 | 0.91394
7.7 | 0.91315
7.75 | 0.91229
7.8 | 0.91156
7.85 | 0.91074
7.9 | 0.91001
7.95 | 0.90913
8 | 0.90823
8.05 | 0.90738
8.1 | 0.90654
8.15 | 0.90566
8.2 | 0.90489
8.25 | 0.90399
8.3 | 0.90328
8.35 | 0.90242
8.4 | 0.90157
8.45 | 0.90072
8.5 | 0.89976
8.55 | 0.89882
8.6 | 0.89795
8.65 | 0.89703
8.7 | 0.89609
8.75 | 0.89524
8.8 | 0.89435
8.85 | 0.89349
8.9 | 0.89266
8.95 | 0.8918
9 | 0.89085
9.05 | 0.88996
9.1 | 0.88905
9.15 | 0.8881
9.2 | 0.88717
9.25 | 0.88622
9.3 | 0.88532
9.35 | 0.88434
9.4 | 0.88341
9.45 | 0.88254
9.5 | 0.88158
9.55 | 0.88067
9.6 | 0.87965
9.65 | 0.87879
9.7 | 0.87786
9.75 | 0.87704
9.8 | 0.87601
9.85 | 0.87509
9.9 | 0.87419
9.95 | 0.87324
10 | 0.87241
10.05 | 0.87149
10.1 | 0.87049
10.15 | 0.86956
10.2 | 0.86862
10.25 | 0.86765
10.3 | 0.86666
10.35 | 0.86573
10.4 | 0.86481
10.45 | 0.86387
10.5 | 0.8629
10.55 | 0.86195
10.6 | 0.86099
10.65 | 0.85996
10.7 | 0.85891
10.75 | 0.85801
10.8 | 0.85705
10.85 | 0.85608
10.9 | 0.85517
10.95 | 0.8541
11 | 0.85319
11.05 | 0.85212
11.1 | 0.85114
11.15 | 0.85018
11.2 | 0.84912
11.25 | 0.84814
11.3 | 0.84709
11.35 | 0.84607
11.4 | 0.8451
11.45 | 0.84408
11.5 | 0.84305
11.55 | 0.8421
11.6 | 0.84107
11.65 | 0.84013
11.7 | 0.83909
11.75 | 0.8381
11.8 | 0.83715
11.85 | 0.83614
11.9 | 0.83514
11.95 | 0.83412
12 | 0.83311
12.05 | 0.83197
12.1 | 0.83099
12.15 | 0.82994
12.2 | 0.82895
12.25 | 0.82789
12.3 | 0.82684
12.35 | 0.82568
12.4 | 0.82473
12.45 | 0.82362
12.5 | 0.82254
12.55 | 0.82147
12.6 | 0.82039
12.65 | 0.81937
12.7 | 0.81824
12.75 | 0.81731
12.8 | 0.81627
12.85 | 0.81526
12.9 | 0.81417
12.95 | 0.81319
13 | 0.81219
13.05 | 0.81107
13.1 | 0.80998
13.15 | 0.80892
13.2 | 0.80794
13.25 | 0.80687
13.3 | 0.80583
13.35 | 0.80479
13.4 | 0.80377
13.45 | 0.80265
13.5 | 0.80158
13.55 | 0.80057
13.6 | 0.79943
13.65 | 0.79835
13.7 | 0.79731
13.75 | 0.79613
13.8 | 0.79507
13.85 | 0.79396
13.9 | 0.79288
13.95 | 0.79184
14 | 0.79085
14.05 | 0.7897
14.1 | 0.78854
14.15 | 0.78738
14.2 | 0.78635
14.25 | 0.78531
14.3 | 0.78416
14.35 | 0.78304
14.4 | 0.78203
14.45 | 0.78098
14.5 | 0.77986
14.55 | 0.77867
14.6 | 0.77753
14.65 | 0.77646
14.7 | 0.7754
14.75 | 0.7743
14.8 | 0.77311
14.85 | 0.77198
14.9 | 0.77086
14.95 | 0.76984
15 | 0.76881
15.05 | 0.76766
15.1 | 0.7665
15.15 | 0.76548
15.2 | 0.76432
15.25 | 0.76313
15.3 | 0.7621
15.35 | 0.76093
15.4 | 0.75987
15.45 | 0.75879
15.5 | 0.75768
15.55 | 0.75653
15.6 | 0.75542
15.65 | 0.75428
15.7 | 0.75306
15.75 | 0.75201
15.8 | 0.75077
15.85 | 0.74963
15.9 | 0.74858
15.95 | 0.74751
16 | 0.74639
16.05 | 0.74537
16.1 | 0.74415
16.15 | 0.74305
16.2 | 0.74186
16.25 | 0.74073
16.3 | 0.73949
16.35 | 0.73835
16.4 | 0.73724
16.45 | 0.73606
16.5 | 0.73497
16.55 | 0.73383
16.6 | 0.73265
16.65 | 0.73146
16.7 | 0.73027
16.75 | 0.72913
16.8 | 0.72795
16.85 | 0.72678
16.9 | 0.72565
16.95 | 0.72463
17 | 0.72344
17.05 | 0.72226
17.1 | 0.72113
17.15 | 0.71998
17.2 | 0.7188
17.25 | 0.71755
17.3 | 0.71634
17.35 | 0.71523
17.4 | 0.7141
17.45 | 0.71298
17.5 | 0.71179
17.55 | 0.71064
17.6 | 0.70953
17.65 | 0.70841
17.7 | 0.7073
17.75 | 0.70615
17.8 | 0.70505
17.85 | 0.70385
17.9 | 0.70262
17.95 | 0.70148
18 | 0.7004
18.05 | 0.69929
18.1 | 0.69808
18.15 | 0.69683
18.2 | 0.69566
18.25 | 0.69447
18.3 | 0.69331
18.35 | 0.69208
18.4 | 0.69092
18.45 | 0.68984
18.5 | 0.68871
18.55 | 0.68758
18.6 | 0.68646
18.65 | 0.68536
18.7 | 0.68416
18.75 | 0.68303
18.8 | 0.68178
18.85 | 0.68067
18.9 | 0.67954
18.95 | 0.67834
19 | 0.67715
19.05 | 0.67592
19.1 | 0.67477
19.15 | 0.67346
19.2 | 0.6723
19.25 | 0.67117
19.3 | 0.67003
19.35 | 0.66899
19.4 | 0.66789
19.45 | 0.66674
19.5 | 0.66552
19.55 | 0.66429
19.6 | 0.66311
19.65 | 0.66191
19.7 | 0.66071
19.75 | 0.65965
19.8 | 0.65845
19.85 | 0.6573
19.9 | 0.6561
19.95 | 0.65487
20 | 0.65371

## F(r,t)
| Anisotropy Function Table | [5 empty] | 137Cs - LDR Bebig CSM40
F(r,q)
r (cm)
| 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 5 | 6 | 7 | 8 | 10 | 12 | 15
0 |  |  |  |  | 0.95752 | 0.95237 | 0.94839 | 0.94646 | 0.94617 | 0.9477 | 0.95025 | 0.94885 | 0.94975 | 0.95405 | 0.95333 | 0.95593
1 |  |  |  |  | 0.96069 | 0.95492 | 0.95201 | 0.94995 | 0.9488 | 0.94954 | 0.95071 | 0.95048 | 0.95145 | 0.95512 | 0.95609 | 0.95891
2 |  |  |  |  | 0.96387 | 0.95746 | 0.95563 | 0.95343 | 0.95142 | 0.95138 | 0.95118 | 0.9521 | 0.95314 | 0.95619 | 0.95886 | 0.9619
3 |  |  |  |  | 0.96276 | 0.95625 | 0.95209 | 0.94831 | 0.94501 | 0.94513 | 0.9459 | 0.94723 | 0.94863 | 0.95108 | 0.95447 | 0.95897
4 |  |  |  |  | 0.9584 | 0.95029 | 0.94483 | 0.94186 | 0.94099 | 0.94272 | 0.94411 | 0.94583 | 0.94848 | 0.95176 | 0.95421 | 0.95843
5 |  |  |  |  | 0.95288 | 0.94439 | 0.94107 | 0.93934 | 0.93989 | 0.94171 | 0.94391 | 0.94554 | 0.94776 | 0.95146 | 0.95517 | 0.95929
6 |  |  |  |  | 0.94623 | 0.94076 | 0.93979 | 0.93959 | 0.9415 | 0.9435 | 0.94594 | 0.94829 | 0.95063 | 0.95374 | 0.95614 | 0.96151
8 |  |  |  | 0.94937 | 0.94579 | 0.94411 | 0.94435 | 0.94464 | 0.94665 | 0.94956 | 0.9513 | 0.95361 | 0.95518 | 0.9584 | 0.96138 | 0.96439
10 |  |  |  | 0.96548 | 0.95262 | 0.95108 | 0.95123 | 0.95209 | 0.95413 | 0.95638 | 0.9584 | 0.95979 | 0.96129 | 0.96399 | 0.96629 | 0.96884
15 |  |  | 1.046 | 0.9862 | 0.97023 | 0.96784 | 0.96778 | 0.96768 | 0.96878 | 0.97045 | 0.9714 | 0.97229 | 0.973 | 0.97505 | 0.97642 | 0.978
20 |  | 1.0443 | 1.0233 | 0.9947 | 0.98034 | 0.97789 | 0.97824 | 0.97775 | 0.9783 | 0.97946 | 0.97974 | 0.98049 | 0.98113 | 0.98183 | 0.98294 | 0.98416
25 |  | 1.0262 | 1.0137 | 0.99784 | 0.9873 | 0.98474 | 0.98445 | 0.98428 | 0.98451 | 0.98527 | 0.98567 | 0.98572 | 0.98617 | 0.98685 | 0.98763 | 0.98769
30 |  | 1.0168 | 1.0091 | 0.99952 | 0.99169 | 0.9896 | 0.98887 | 0.98874 | 0.98888 | 0.98951 | 0.98952 | 0.98939 | 0.98981 | 0.99011 | 0.99055 | 0.99127
40 | 1.0335 | 1.0081 | 1.0043 | 1.0001 | 0.99609 | 0.99475 | 0.99413 | 0.99379 | 0.99404 | 0.99416 | 0.99439 | 0.99467 | 0.99498 | 0.99478 | 0.99515 | 0.99499
50 | 1.0156 | 1.0038 | 1.0021 | 1.0004 | 0.99825 | 0.9973 | 0.99702 | 0.9968 | 0.99693 | 0.99701 | 0.99723 | 0.99718 | 0.99716 | 0.99732 | 0.99703 | 0.99711
60 | 1.0074 | 1.0018 | 1.0012 | 1.0003 | 0.99926 | 0.99869 | 0.9988 | 0.99829 | 0.9985 | 0.99853 | 0.99875 | 0.99828 | 0.99864 | 0.99891 | 0.99862 | 0.99875
70 | 1.003 | 1.0005 | 1.0004 | 1.0003 | 0.99979 | 0.99959 | 0.99952 | 0.99959 | 0.9996 | 0.99945 | 0.99941 | 0.99951 | 0.99918 | 0.9993 | 0.99934 | 0.99926
80 | 1.001 | 1 | 1.0003 | 1.0002 | 1.0003 | 0.99996 | 1.0001 | 0.9996 | 0.9997 | 1.0003 | 1.0001 | 0.99996 | 0.99976 | 1.0001 | 1 | 0.99988
90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
100 | 1.0008 | 1.0001 | 0.99988 | 1.0002 | 1.0002 | 0.99993 | 0.99989 | 0.99981 | 1 | 1.0003 | 1.0001 | 1 | 0.99982 | 0.99993 | 0.99991 | 0.99986
110 | 1.0028 | 1.0005 | 1.0002 | 1.0001 | 1 | 0.99945 | 0.99937 | 0.99921 | 0.99929 | 0.99943 | 0.99943 | 0.99941 | 0.99936 | 0.9993 | 0.99935 | 0.99941
120 | 1.007 | 1.0018 | 1.0007 | 1.0004 | 0.99911 | 0.99902 | 0.99831 | 0.99825 | 0.99816 | 0.99853 | 0.99885 | 0.99884 | 0.99883 | 0.99907 | 0.99885 | 0.99887
130 | 1.0154 | 1.0037 | 1.0019 | 1.0006 | 0.99811 | 0.9974 | 0.99732 | 0.99669 | 0.99687 | 0.99741 | 0.9974 | 0.99701 | 0.99696 | 0.9975 | 0.99727 | 0.99753
140 | 1.033 | 1.008 | 1.0038 | 1.0002 | 0.99586 | 0.99425 | 0.99446 | 0.99402 | 0.99439 | 0.99498 | 0.99487 | 0.99476 | 0.99492 | 0.99489 | 0.99521 | 0.99502
150 |  | 1.0168 | 1.0088 | 0.99958 | 0.99163 | 0.98998 | 0.9896 | 0.98904 | 0.98885 | 0.98982 | 0.9898 | 0.99013 | 0.99067 | 0.99084 | 0.99118 | 0.99141
155 |  | 1.0263 | 1.0137 | 0.99803 | 0.98811 | 0.98588 | 0.98504 | 0.98517 | 0.98539 | 0.98575 | 0.98598 | 0.98589 | 0.98652 | 0.98719 | 0.98768 | 0.98847
160 |  | 1.0442 | 1.0235 | 0.99447 | 0.9816 | 0.97914 | 0.97853 | 0.97835 | 0.979 | 0.97976 | 0.98037 | 0.98083 | 0.98141 | 0.98272 | 0.98311 | 0.98477
165 |  |  | 1.0461 | 0.98645 | 0.97209 | 0.96912 | 0.96906 | 0.96858 | 0.9694 | 0.97106 | 0.97164 | 0.97239 | 0.97397 | 0.97495 | 0.97693 | 0.97867
170 |  |  |  | 0.96742 | 0.95509 | 0.95307 | 0.95279 | 0.95264 | 0.95437 | 0.95693 | 0.95805 | 0.96015 | 0.96184 | 0.96492 | 0.9666 | 0.96946
172 |  |  |  | 0.95612 | 0.94666 | 0.94467 | 0.94439 | 0.94472 | 0.94676 | 0.94929 | 0.95178 | 0.95382 | 0.95559 | 0.9586 | 0.96169 | 0.96491
174 |  |  |  |  | 0.94436 | 0.93983 | 0.93913 | 0.93898 | 0.94025 | 0.94233 | 0.94487 | 0.94698 | 0.94931 | 0.95244 | 0.95629 | 0.96108
175 |  |  |  |  | 0.95 | 0.94177 | 0.93817 | 0.93702 | 0.93813 | 0.94046 | 0.94272 | 0.94517 | 0.94735 | 0.95107 | 0.95518 | 0.9594
176 |  |  |  |  | 0.95424 | 0.94574 | 0.94203 | 0.94043 | 0.93869 | 0.93987 | 0.94252 | 0.94462 | 0.94613 | 0.9498 | 0.95369 | 0.95794
177 |  |  |  |  | 0.95691 | 0.95162 | 0.94757 | 0.94461 | 0.94283 | 0.94362 | 0.945 | 0.94586 | 0.94703 | 0.9505 | 0.95289 | 0.95731
178 |  |  |  |  | 0.95988 | 0.95292 | 0.94949 | 0.94787 | 0.94633 | 0.94615 | 0.94649 | 0.94837 | 0.94947 | 0.95262 | 0.95418 | 0.95878
179 |  |  |  |  | 0.95887 | 0.95399 | 0.95161 | 0.95059 | 0.94829 | 0.94896 | 0.9495 | 0.95058 | 0.95045 | 0.95416 | 0.95507 | 0.95771
180 |  |  |  |  | 0.95787 | 0.95506 | 0.95374 | 0.95332 | 0.95024 | 0.95178 | 0.9525 | 0.95278 | 0.95142 | 0.95571 | 0.95596 | 0.95664

## F(r,t) - plot
Anisotropy Function Table | 137Cs - LDR Bebig CSM40

## D(y,z)
Absolute dose rate table: D(y,z) (cGy/h U)
137Cs - LDR Bebig CSM40
cGy/(h U)
Along, | [7 empty] | Away, | y (cm)
z (cm) | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 5 | 6 | 8 | 10
-10 | 0.0093398 | 0.0092827 | 0.0092257 | 0.0091884 | 0.0091604 | 0.0091107 | 0.0090262 | 0.008849 | 0.0086449 | 0.0081024 | 0.0074724 | 0.006803 | 0.0054988 | 0.0043536
-8 | 0.01521 | 0.015105 | 0.015 | 0.014931 | 0.014872 | 0.014748 | 0.014448 | 0.014022 | 0.013512 | 0.012282 | 0.010952 | 0.0096416 | 0.0073193 | 0.0055064
-6 | 0.028089 | 0.027841 | 0.027593 | 0.027498 | 0.027397 | 0.026771 | 0.025698 | 0.024365 | 0.022841 | 0.019633 | 0.016552 | 0.013843 | 0.0096561 | 0.0068337
-5 | 0.041349 | 0.040918 | 0.040488 | 0.040353 | 0.040006 | 0.038545 | 0.036289 | 0.033635 | 0.030812 | 0.025283 | 0.020468 | 0.01654 | 0.010975 | 0.00752
-4 | 0.066162 | 0.065509 | 0.064857 | 0.064376 | 0.063191 | 0.059162 | 0.053857 | 0.048161 | 0.042592 | 0.032815 | 0.02525 | 0.019611 | 0.012332 | 0.008176
-3 | 0.12255 | 0.12128 | 0.12 | 0.11712 | 0.11219 | 0.099046 | 0.084665 | 0.071364 | 0.059817 | 0.042346 | 0.03066 | 0.02283 | 0.013617 | 0.0087696
-2.5 | 0.18479 | 0.18148 | 0.17817 | 0.17059 | 0.15931 | 0.13312 | 0.10824 | 0.087342 | 0.070759 | 0.047651 | 0.03345 | 0.024367 | 0.014188 | 0.0090173
-2 | 0.31605 | 0.3038 | 0.29154 | 0.2676 | 0.23873 | 0.18304 | 0.13892 | 0.10648 | 0.082931 | 0.053049 | 0.036092 | 0.025802 | 0.014694 | 0.009235
-1.5 | 0.7034 | 0.62982 | 0.55624 | 0.46211 | 0.37717 | 0.25344 | 0.17639 | 0.12753 | 0.095427 | 0.058104 | 0.038448 | 0.027013 | 0.015111 | 0.009415
-1 |  | 2.1699 | 1.289 | 0.8468 | 0.59764 | 0.33989 | 0.21599 | 0.14781 | 0.10674 | 0.062276 | 0.040293 | 0.027959 | 0.015422 | 0.0095418
-0.75 |  | 4.7153 | 1.9036 | 1.0966 | 0.7226 | 0.38188 | 0.23362 | 0.15632 | 0.11125 | 0.063878 | 0.040988 | 0.028291 | 0.015525 | 0.0095869
-0.5 |  | 6.3595 | 2.4169 | 1.3136 | 0.83067 | 0.41687 | 0.24775 | 0.16293 | 0.11471 | 0.065054 | 0.041481 | 0.028551 | 0.015609 | 0.0096193
-0.25 |  | 6.9005 | 2.6933 | 1.4507 | 0.90175 | 0.43976 | 0.25687 | 0.16714 | 0.11687 | 0.065788 | 0.041793 | 0.028705 | 0.015656 | 0.0096395
0 |  |  | 2.7765 | 1.4962 | 0.92633 | 0.44778 | 0.26003 | 0.16856 | 0.11762 | 0.066048 | 0.041905 | 0.028745 | 0.015668 | 0.0096441
0.25 |  | 6.9008 | 2.6934 | 1.4509 | 0.90201 | 0.4399 | 0.25687 | 0.16711 | 0.11685 | 0.065767 | 0.041784 | 0.0287 | 0.015654 | 0.0096402
0.5 |  | 6.3589 | 2.4169 | 1.3138 | 0.83047 | 0.41681 | 0.2477 | 0.16295 | 0.11469 | 0.06505 | 0.041497 | 0.028542 | 0.015602 | 0.0096222
0.75 |  | 4.7155 | 1.9042 | 1.0967 | 0.72278 | 0.38196 | 0.23367 | 0.15633 | 0.11127 | 0.06386 | 0.040992 | 0.028302 | 0.015529 | 0.0095873
1 |  | 2.1683 | 1.2892 | 0.84674 | 0.59769 | 0.33994 | 0.216 | 0.14785 | 0.10671 | 0.062269 | 0.040305 | 0.027958 | 0.015412 | 0.0095406
1.5 | 0.70245 | 0.62877 | 0.55509 | 0.46188 | 0.3772 | 0.25341 | 0.1764 | 0.12757 | 0.09549 | 0.058071 | 0.038442 | 0.027016 | 0.015104 | 0.0094134
2 | 0.31636 | 0.30382 | 0.29127 | 0.26732 | 0.23857 | 0.18304 | 0.13887 | 0.10642 | 0.08295 | 0.053038 | 0.036104 | 0.0258 | 0.014693 | 0.009237
2.5 | 0.18526 | 0.18166 | 0.17807 | 0.17037 | 0.15914 | 0.13309 | 0.10815 | 0.08735 | 0.070766 | 0.04766 | 0.033458 | 0.02438 | 0.014184 | 0.0090196
3 | 0.12309 | 0.12148 | 0.11987 | 0.11706 | 0.11215 | 0.098951 | 0.084641 | 0.071304 | 0.059798 | 0.042343 | 0.030663 | 0.022835 | 0.013614 | 0.0087661
4 | 0.066357 | 0.065637 | 0.064917 | 0.064352 | 0.063117 | 0.059111 | 0.053843 | 0.04815 | 0.042564 | 0.032808 | 0.025236 | 0.019616 | 0.012334 | 0.0081762
5 | 0.041412 | 0.040966 | 0.04052 | 0.040327 | 0.039993 | 0.038512 | 0.036291 | 0.033602 | 0.030805 | 0.02528 | 0.020471 | 0.016544 | 0.010981 | 0.0075194
6 | 0.028299 | 0.027958 | 0.027617 | 0.027497 | 0.027376 | 0.02676 | 0.025717 | 0.024363 | 0.022821 | 0.019618 | 0.016556 | 0.013836 | 0.0096496 | 0.0068362
8 | 0.015292 | 0.015153 | 0.015014 | 0.014919 | 0.014879 | 0.014733 | 0.014439 | 0.014024 | 0.013509 | 0.01228 | 0.010951 | 0.0096427 | 0.0073195 | 0.0055084
10 | 0.0094136 | 0.0093242 | 0.0092349 | 0.0092008 | 0.0091774 | 0.0090988 | 0.0090102 | 0.0088529 | 0.0086351 | 0.0080962 | 0.0074799 | 0.0068052 | 0.0054963 | 0.0043527