## General Information
| General Information | I-125 - Amersham Health model 6702
Source Information
Source name: | Amersham Health model 6702
Radionuclide: | I-125
Model Number: | 6702
Distributed by: | Amersham Health
Manufactured by:
References
1) | "Update of AAPM Task Group No. 43 Report: A revised AAPM protocol for brachytherapy dose calculations"
Med Phys 31 (3), March 2004, 633-674
M.J. Rivard, B.M. Coursey, L.A. DeWerd, W.F. Hanson, M.S. Huq, G.S. Ibbott, M.G. Mitch, R. Nath,J.F. Williamson
Calculation details
Method: | Monte Carlo and TLD averaged
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised
2006-01-26

## Source design and materials
Source design and materials | 125I - Amersham Health model 6702

## Geometry Function
Geometry Function |  |  |  | 125I - Amersham Health model 6702
Active source length,
| L = | 3 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | 125I - Amersham Health model 6702
L = | 1.036 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | 125I - Amersham Health model 6702
F(r,q)
r (cm)
q (º) | 0.5 | 1 | 2 | 3 | 4 | 5
0 | 0.385 | 0.42 | 0.493 | 0.533 | 0.569 | 0.589
5 | 0.413 | 0.472 | 0.546 | 0.586 | 0.613 | 0.631
10 | 0.531 | 0.584 | 0.63 | 0.66 | 0.681 | 0.697
15 | 0.7 | 0.7 | 0.719 | 0.738 | 0.749 | 0.758
20 | 0.788 | 0.789 | 0.793 | 0.805 | 0.81 | 0.814
30 | 0.892 | 0.888 | 0.888 | 0.891 | 0.892 | 0.892
40 | 0.949 | 0.948 | 0.944 | 0.944 | 0.944 | 0.944
50 | 0.977 | 0.973 | 0.967 | 0.967 | 0.967 | 0.967
60 | 0.989 | 0.985 | 0.983 | 0.983 | 0.983 | 0.983
70 | 0.996 | 0.992 | 0.99 | 0.99 | 0.99 | 0.99
80 | 1 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998

## Radial Dose Function
Radial Dose Function |  |  | 125I - Amersham Health model 6702
Line source aproximation,
L = | 3 | mm | Point source aproximation
r (cm) | g(r) |  | r (cm) | g(r)
0.1 | 1.02 |  | 0.1 | 0.673
0.15 | 1.022 |  | 0.15 | 0.809
0.25 | 1.024 |  | 0.25 | 0.929
0.5 | 1.03 |  | 0.5 | 1.008
0.75 | 1.02 |  | 0.75 | 1.014
1 | 1 |  | 1 | 1
1.5 | 0.935 |  | 1.5 | 0.939
2 | 0.861 |  | 2 | 0.866
3 | 0.697 |  | 3 | 0.702
4 | 0.553 |  | 4 | 0.557
5 | 0.425 |  | 5 | 0.428
6 | 0.322 |  | 6 | 0.324
7 | 0.241 |  | 7 | 0.243
8 | 0.179 |  | 8 | 0.18
9 | 0.134 |  | 9 | 0.135
10 | 0.0979 |  | 10 | 0.0986

## Anisotropy Factor
Anisotropy Factor |  |  | 125I - Amersham Health model 6702
| r (cm) | fan(r)
| 0.5 | 0.986
| 1 | 0.96
| 2 | 0.952
| 3 | 0.951
| 4 | 0.954
| 5 | 0.954

## QA
Dose rates for QA |  |  | I-125 - Amersham Health model 6702
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources using
the 1-D formalism of Eq. (11) with interpolation for gL(r) and Øan(r).
| r (cm) | fan(r)
| 0.5 | 4.119
| 1 | 0.995
| 1.5 | 0.413
| 2 | 0.213
| 3 | 0.0768
| 4 | 0.0344
| 5 | 0.0169
| 6 | 0.0089
| 7 | 0.0049