## General Information
| General Information | I-125 - Amersham model 6733
Source Information
Source name: | Amersham model 6733
Radionuclide: | I-125
Model Number: | 6733
Distributed by: | Amersham Health
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | I-125 - Amersham model 6733

## Geometry Function
Geometry Function |  |  | I-125 - Amersham model 6733
Source length for Line source approximation:
L= | 3 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | I-125 - Amersham model 6733
L = | 0.98 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | I-125 - Amersham model 6733
F(r,q)
r (cm)
q (º) | 1 | 2 | 3 | 4 | 5 | 6 | 7
0 | 0.305 | 0.397 | 0.451 | 0.502 | 0.533 | 0.551 | 0.565
5 | 0.386 | 0.468 | 0.51 | 0.557 | 0.586 | 0.595 | 0.611
10 | 0.507 | 0.57 | 0.609 | 0.634 | 0.66 | 0.669 | 0.685
15 | 0.621 | 0.663 | 0.68 | 0.712 | 0.717 | 0.726 | 0.719
20 | 0.714 | 0.738 | 0.743 | 0.774 | 0.769 | 0.779 | 0.785
30 | 0.848 | 0.851 | 0.849 | 0.873 | 0.859 | 0.86 | 0.88
40 | 0.944 | 0.933 | 0.918 | 0.932 | 0.921 | 0.912 | 0.924
50 | 0.999 | 0.985 | 0.969 | 0.983 | 0.953 | 0.965 | 0.949
60 | 1.029 | 1.015 | 0.995 | 1.012 | 0.985 | 1.003 | 0.982
70 | 1.038 | 1.033 | 1.015 | 1.022 | 1.001 | 0.994 | 1.019
80 | 1.026 | 1.034 | 1.014 | 1.026 | 1.009 | 0.999 | 1
90 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | I-125 - Amersham model 6733
Line source approximation,
L = | 3 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gP(r)
0.1 | 1.05 |  | 0.1 | 0.693
0.15 | 1.076 |  | 0.15 | 0.851
0.25 | 1.085 |  | 0.25 | 0.985
0.5 | 1.069 |  | 0.5 | 1.046
0.75 | 1.045 |  | 0.75 | 1.039
1 | 1 |  | 1 | 1
1.5 | 0.912 |  | 1.5 | 0.916
2 | 0.821 |  | 2 | 0.826
3 | 0.656 |  | 3 | 0.66
4 | 0.495 |  | 4 | 0.498
5 | 0.379 |  | 5 | 0.382
6 | 0.285 |  | 6 | 0.287
7 | 0.214 |  | 7 | 0.216
8 | 0.155 |  | 8 | 0.156
9 | 0.119 |  | 9 | 0.12
10 | 0.084 |  | 10 | 0.0846

## Anisotropy Factor
Anisotropy Factor |  |  | I-125 - Amersham model 6733
| r (cm) | fan(r)
| 0.1 | 1.173
| 0.15 | 1.246
| 0.25 | 1.112
| 0.5 | 0.996
| 0.75 | 0.974
| 1 | 0.967
| 2 | 0.964
| 3 | 0.953
| 4 | 0.966
| 5 | 0.953
| 6 | 0.948
| 7 | 0.955

## QA
Dose rates for QA |  |  | I-125 - Amersham model 6733
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 1.173
| 0.15 | 1.246
| 0.25 | 1.112
| 0.5 | 0.996
| 0.75 | 0.974
| 1 | 0.967
| 2 | 0.964
| 3 | 0.953
| 4 | 0.966
| 5 | 0.953
| 6 | 0.948
| 7 | 0.955