## General Information
| General Information | I-125 - Implant Sciences 3500
Source Information
Source name: | Implant Sciences 3500
Radionuclide: | I-125
Model Number:
Distributed by:
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | I-125 - Implant Sciences 3500

## Geometry Function
Geometry Function |  |  | I-125 - Implant Sciences 3500
Source length for Line source approximation:
L = | 3.76 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | I-125 - Implant Sciences 3500
L = | 1.014 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | I-125 - Implant Sciences 3500
F(r,q)
r (cm)
q (º) | 0.25 | 0.5 | 1 | 2 | 5 | 10
0 | 0.494 | 0.61 | 0.58 | 0.652 | 0.69 | 0.709
10 | 0.574 | 0.513 | 0.561 | 0.626 | 0.7 | 0.742
20 | 0.785 | 0.679 | 0.705 | 0.743 | 0.789 | 0.815
30 | 0.899 | 0.808 | 0.813 | 0.83 | 0.854 | 0.872
40 | 0.943 | 0.892 | 0.885 | 0.893 | 0.905 | 0.912
50 | 0.967 | 0.944 | 0.933 | 0.934 | 0.941 | 0.947
60 | 0.986 | 0.974 | 0.967 | 0.967 | 0.968 | 0.972
70 | 0.995 | 0.99 | 0.987 | 0.987 | 0.986 | 0.99
80 | 1 | 0.997 | 0.997 | 0.997 | 0.996 | 0.997
90 | 1 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | I-125 - Implant Sciences 3500
Line source approximation,
L = | 3.76 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gP(r)
0.1 | 0.997 |  | 0.1 | 0.576
0.15 | 1.011 |  | 0.15 | 0.732
0.25 | 1.021 |  | 0.25 | 0.886
0.5 | 1.03 |  | 0.5 | 0.997
0.75 | 1.026 |  | 0.75 | 1.017
1 | 1 |  | 1 | 1
1.5 | 0.932 |  | 1.5 | 0.938
2 | 0.854 |  | 2 | 0.862
3 | 0.681 |  | 3 | 0.688
4 | 0.532 |  | 4 | 0.538
5 | 0.407 |  | 5 | 0.412
6 | 0.308 |  | 6 | 0.312
7 | 0.23 |  | 7 | 0.233
8 | 0.171 |  | 8 | 0.173
9 | 0.127 |  | 9 | 0.129
10 | 0.0936 |  | 10 | 0.0947

## Anisotropy Factor
Anisotropy Factor |  |  | I-125 - Implant Sciences 3500
| r (cm) | fan(r)
| 0.1 | 1.129
| 0.15 | 1.268
| 0.25 | 1.164
| 0.5 | 0.973
| 0.75 | 0.942
| 1 | 0.933
| 2 | 0.931
| 5 | 0.938
| 10 | 0.948

## QA
Dose rates for QA |  |  | I-125 - Implant Sciences 3500
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 65.6
| 0.15 | 41.7
| 0.25 | 16.7
| 0.5 | 3.93
| 0.75 | 1.73
| 1 | 0.946
| 1.5 | 0.394
| 2 | 0.203
| 3 | 0.0724
| 4 | 0.0319
| 5 | 0.0157
| 6 | 0.00824
| 7 | 0.00454
| 8 | 0.00259
| 9 | 0.00152
| 10 | 0.00091