## General information
General Information
Source Information
Source name: | LDR wire
Radionuclide: | Ir-192
Model Number: | Ir-192 LDR 0.3 mm diameter and 7 cm long
Distributed by: | Amersham
Manufactured by:
References
1) | "Dosimetry of 192Ir wires for LDR interstitial brachytherapy following the AAPM TG-43 dosimetric formalism"
Med Phys 28(2001),156-166
P. Karaiskos, P. Papagiannis, A. Angelopoulos, L.Sakelliou, D. Baltas, P. Sandilos, L. Vlachos
Calculation details
Method: | Monte Carlo
Code:
Version:
Phantom material: | Water
Phantom size: | Sphere 30 cm in diameter
Cutoff energy: | 10 keV
Cross-section:
Electron tracking: | No
Revised
2006-01-31

## Source design and materials
Source design and materials | Ir-192 LDR 0.3 mm diameter and 7 cm long
The central core is 0.1 mm in diameter, composed of about 25% Ir and 75% Pt, encased in a 0.1 mm Pt sheath
Coordinate system is situated in the source center.

## Geometry Function
Geometry Function |  |  |  | Ir-192 LDR 0.3 mm diameter and 7 cm long
Geometry factor G(r,q) for L = 5 cm.
|  |  |  | IMPORTANT: | To use these data to obtain geometry factors for L = 7 cm, see table I and Section II.B in Karaiskos et al.
Radial distance r (cm)
Angle (º) | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 5 | 6 | 7 | 10
1 | 205.4 | 179.9 | 119.8 | 89.8 | 71.8 | 59.8 | 51.3 | 44.8 | 39.8 | 35.811 | 23.751 | 17.557 | 7.2 | 0.363 | 0.103 | 0.053 | 0.034 | 0.023 | 0.011
2 | 102.5 | 89.9 | 59.9 | 44.8 | 35.8 | 29.8 | 25.6 | 22.3 | 19.8 | 17.813 | 11.753 | 8.561 | 3.601 | 0.359 | 0.102 | 0.053 | 0.034 | 0.023 | 0.011
3 | 68.3 | 59.9 | 39.9 | 29.9 | 23.8 | 19.8 | 17 | 14.8 | 13.2 | 11.815 | 7.754 | 5.566 | 2.401 | 0.354 | 0.102 | 0.053 | 0.034 | 0.023 | 0.011
4 | 51.1 | 44.9 | 29.9 | 22.4 | 17.9 | 14.8 | 12.7 | 11.1 | 9.83 | 8.817 | 5.756 | 4.073 | 1.801 | 0.347 | 0.102 | 0.053 | 0.034 | 0.023 | 0.011
5 | 40.9 | 35.9 | 23.9 | 17.9 | 14.3 | 11.6 | 10.1 | 8.833 | 7.83 | 7.019 | 4.558 | 3.181 | 1.442 | 0.339 | 0.101 | 0.053 | 0.034 | 0.023 | 0.011
7 | 29.1 | 25.6 | 17 | 12.7 | 10.2 | 8.42 | 7.19 | 6.267 | 5.53 | 4.966 | 3.191 | 2.171 | 1.031 | 0.321 | 0.101 | 0.053 | 0.033 | 0.023 | 0.011
10 | 20.3 | 17.9 | 11.9 | 8.88 | 7.07 | 5.86 | 5 | 4.346 | 3.84 | 3.43 | 2.171 | 1.431 | 0.724 | 0.291 | 0.099 | 0.052 | 0.033 | 0.023 | 0.011
15 | 13.5 | 12 | 7.93 | 5.91 | 4.69 | 3.88 | 3.3 | 2.858 | 2.52 | 2.241 | 1.386 | 0.883 | 0.486 | 0.246 | 0.094 | 0.051 | 0.033 | 0.023 | 0.011
20 | 10.2 | 9.03 | 5.96 | 4.43 | 3.51 | 2.89 | 2.45 | 2.122 | 1.86 | 1.653 | 1.002 | 0.629 | 0.367 | 0.21 | 0.089 | 0.05 | 0.032 | 0.023 | 0.011
25 | 8.16 | 7.27 | 4.79 | 3.55 | 2.81 | 2.31 | 1.95 | 1.685 | 1.48 | 1.306 | 0.78 | 0.487 | 0.297 | 0.183 | 0.085 | 0.049 | 0.032 | 0.023 | 0.011
30 | 6.85 | 6.12 | 4.03 | 2.98 | 2.35 | 1.93 | 1.63 | 1.4 | 1.22 | 1.079 | 0.637 | 0.399 | 0.251 | 0.163 | 0.08 | 0.047 | 0.031 | 0.023 | 0.01
35 | 5.93 | 5.32 | 3.49 | 2.58 | 2.03 | 1.66 | 1.4 | 1.2 | 1.05 | 0.921 | 0.539 | 0.339 | 0.219 | 0.147 | 0.076 | 0.046 | 0.03 | 0.022 | 0.01
40 | 5.26 | 4.73 | 3.1 | 2.28 | 1.79 | 1.47 | 1.23 | 1.055 | 0.917 | 0.807 | 0.469 | 0.297 | 0.195 | 0.134 | 0.072 | 0.044 | 0.03 | 0.022 | 0.01
45 | 4.75 | 4.28 | 2.8 | 2.06 | 1.62 | 1.32 | 1.11 | 0.946 | 0.821 | 0.721 | 0.418 | 0.266 | 0.178 | 0.124 | 0.068 | 0.043 | 0.029 | 0.022 | 0.01
50 | 4.36 | 3.94 | 2.57 | 1.89 | 1.48 | 1.2 | 1.01 | 0.862 | 0.747 | 0.656 | 0.379 | 0.242 | 0.164 | 0.116 | 0.066 | 0.042 | 0.029 | 0.021 | 0.01
60 | 3.82 | 3.47 | 2.26 | 1.65 | 1.29 | 1.05 | 0.88 | 0.747 | 0.647 | 0.566 | 0.327 | 0.211 | 0.145 | 0.105 | 0.061 | 0.04 | 0.028 | 0.021 | 0.01
70 | 3.5 | 3.18 | 2.07 | 1.51 | 1.18 | 0.96 | 0.8 | 0.679 | 0.587 | 0.513 | 0.296 | 0.192 | 0.134 | 0.098 | 0.058 | 0.038 | 0.027 | 0.02 | 0.01

## Dose Rate Constant
Dose-Rate Constant |  |  | Ir-192 LDR 0.3 mm diameter and 7 cm long
L = | 0.403 | cGy/(h·U)

## Radial Dose Function
Radial-Dose Function Table |  | Ir-192 LDR 0.3 mm diameter and 7 cm long
r (cm) | g( r )
0.2 | 0.949
0.3 | 0.964
0.5 | 0.98
0.7 | 0.989
1 | 1
1.5 | 1.011
2 | 1.017
3 | 1.02
4 | 1.017
5 | 1.01
6 | 0.997
7 | 0.981
8 | 0.956
9 | 0.93
10 | 0.897

## Anisotropy Function
| Anisotropy Function Table |  |  |  |  | Ir-192 LDR 0.3 mm diameter and 7 cm long
F(r,q)
r (cm)
q (º) | 0.5 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 10
0.5 |  |  |  |  | 0.48 | 0.52 | 0.417 | 0.46 | 0.489 | 0.503
2.5 | 0.928 | 0.927 | 0.915 | 0.928 | 0.703 | 0.596 | 0.6 | 0.593 | 0.6 | 0.624
5.5 | 0.951 | 0.946 | 0.952 | 0.953 | 0.84 | 0.757 | 0.747 | 0.748 | 0.749 | 0.767
7.5 | 0.957 | 0.951 | 0.964 | 0.959 | 0.881 | 0.828 | 0.808 | 0.811 | 0.812 | 0.824
10.5 | 0.962 | 0.956 | 0.973 | 0.964 | 0.924 | 0.872 | 0.863 | 0.856 | 0.862 | 0.872
15.5 | 0.968 | 0.961 | 0.982 | 0.97 | 0.952 | 0.922 | 0.922 | 0.92 | 0.914 | 0.915
20.5 | 0.972 | 0.966 | 0.986 | 0.974 | 0.96 | 0.956 | 0.952 | 0.949 | 0.951 | 0.951
25.5 | 0.975 | 0.969 | 0.99 | 0.977 | 0.968 | 0.971 | 0.968 | 0.97 | 0.963 | 0.965
30.5 | 0.978 | 0.972 | 0.992 | 0.979 | 0.974 | 0.975 | 0.978 | 0.979 | 0.971 | 0.972
35.5 | 0.98 | 0.976 | 0.994 | 0.982 | 0.978 | 0.977 | 0.98 | 0.981 | 0.983 | 0.985
40.5 | 0.983 | 0.979 | 0.996 | 0.985 | 0.982 | 0.981 | 0.983 | 0.986 | 0.988 | 0.99
45.5 | 0.985 | 0.982 | 0.998 | 0.987 | 0.984 | 0.987 | 0.99 | 0.992 | 0.991 | 0.993
50.5 | 0.988 | 0.985 | 1 | 0.989 | 0.987 | 0.991 | 0.992 | 0.993 | 0.993 | 0.995
55.5 | 0.99 | 0.988 | 0.997 | 0.991 | 0.993 | 0.993 | 0.995 | 0.995 | 0.996 | 0.997
60.5 | 0.993 | 0.991 | 0.998 | 0.994 | 0.998 | 0.998 | 0.997 | 0.996 | 0.998 | 0.999
65.5 | 0.993 | 0.994 | 1 | 0.996 | 0.995 | 0.997 | 1.001 | 0.996 | 0.999 | 0.999
70.5 | 0.995 | 0.997 | 0.999 | 0.998 | 0.997 | 0.999 | 1 | 1.001 | 1.001 | 1.001
75.5 | 0.997 | 0.998 | 1 | 1 | 1 | 0.998 | 1 | 0.999 | 1 | 1
80.5 | 0.998 | 1 | 1 | 1 | 0.998 | 0.999 | 0.999 | 1 | 0.999 | 0.998
85.5 | 1 | 1 | 1 | 0.998 | 0.999 | 0.998 | 1 | 1.001 | 1 | 1
90 | bar: 0.5 to r (cm) 10