## M-19
| TG-43 CONSENSUS | [5 empty] | SPEC inc. Co. HDR 192Ir M-19
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.114 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL(r) |  |  |  | F(r,q) | [9 empty] | QA Along-away
| L = | 0.35 | cm |  |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL( r) |  | Theta/deg | 0 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 5 | 7.5 | 10 | z / cm | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 0.99261 |  | 0 | 0.679 | 0.679 | 0.6606 | 0.6238 | 0.6331 | 0.6545 | 0.676 | 0.7006 | 0.7408 | 0.7743 | 7 | 0.016562 | 0.016835 | 0.017163 | 0.017409 | 0.017626 | 0.017857 | 0.01783 | 0.017008 | 0.015544 | 0.013827 | 0.012043 | 0.010457
| 0.25 | 0.99261 |  | 1 | 0.68125 | 0.68125 | 0.6619 | 0.6232 | 0.6359 | 0.6566 | 0.6812 | 0.7039 | 0.7458 | 0.7768 | 6 | 0.022276 | 0.022741 | 0.023309 | 0.023677 | 0.024069 | 0.024343 | 0.024156 | 0.02247 | 0.01994 | 0.01711 | 0.014537 | 0.012276
| 0.5 | 0.99515 |  | 2 | 0.67935 | 0.67935 | 0.6608 | 0.6237 | 0.6432 | 0.6667 | 0.6924 | 0.7119 | 0.7537 | 0.7828 | 5 | 0.031536 | 0.032459 | 0.03339 | 0.034103 | 0.034683 | 0.034871 | 0.034122 | 0.030497 | 0.025836 | 0.021365 | 0.017511 | 0.014333
| 0.75 | 0.998 |  | 3 | 0.6741 | 0.6741 | 0.6604 | 0.633 | 0.6544 | 0.6775 | 0.7026 | 0.7246 | 0.7621 | 0.7903 | 4 | 0.047795 | 0.049922 | 0.051826 | 0.0534 | 0.054132 | 0.053699 | 0.050837 | 0.042526 | 0.033772 | 0.026559 | 0.020884 | 0.016566
| 1 | 1 |  | 5 | 0.67695 | 0.67695 | 0.671 | 0.6591 | 0.681 | 0.7026 | 0.7238 | 0.7437 | 0.7771 | 0.8013 | 3 | 0.08275 | 0.087866 | 0.092432 | 0.09512 | 0.095477 | 0.090184 | 0.080824 | 0.060387 | 0.044088 | 0.03248 | 0.024459 | 0.018811
| 1.5 | 1.0014 |  | 7 | 0.6995 | 0.6995 | 0.6946 | 0.6848 | 0.7043 | 0.7226 | 0.7428 | 0.7586 | 0.7926 | 0.816 | 2 | 0.18029 | 0.19793 | 0.21095 | 0.21149 | 0.20183 | 0.16956 | 0.13609 | 0.085391 | 0.055789 | 0.038498 | 0.027793 | 0.020768
| 2 | 1.0045 |  | 10 | 0.73505 | 0.73505 | 0.7318 | 0.7253 | 0.7402 | 0.7571 | 0.775 | 0.7893 | 0.8178 | 0.8363 | 1.5 | 0.31909 | 0.3585 | 0.37614 | 0.35786 | 0.32146 | 0.24125 | 0.17642 | 0.09941 | 0.061575 | 0.041099 | 0.029149 | 0.021557
| 3 | 1.0078 |  | 12 | 0.7594 | 0.7594 | 0.7563 | 0.7501 | 0.765 | 0.7781 | 0.7952 | 0.8069 | 0.8309 | 0.8484 | 1 | 0.72413 | 0.84391 | 0.80841 | 0.67776 | 0.54144 | 0.33961 | 0.22312 | 0.1126 | 0.066337 | 0.043254 | 0.030253 | 0.022143
| 4 | 1.0034 |  | 15 | 0.79155 | 0.79155 | 0.7901 | 0.7872 | 0.7974 | 0.8085 | 0.8213 | 0.8305 | 0.8528 | 0.8667 | 0.5 | 3.3721 | 3.4151 | 2.1913 | 1.355 | 0.88509 | 0.44694 | 0.26445 | 0.12216 | 0.069642 | 0.044631 | 0.030889 | 0.022481
| 5 | 0.9988 |  | 20 | 0.882 | 0.882 | 0.8405 | 0.8348 | 0.8435 | 0.8521 | 0.8628 | 0.8698 | 0.8843 | 0.8937 | 0 | 390043975.2914 | 15.5922 | 4.3087 | 1.9614 | 1.114 | 0.49857 | 0.28188 | 0.12586 | 0.070522 | 0.044939 | 0.031048 | 0.022583
| 6 | 0.99356 |  | 25 | 0.9131 | 0.9131 | 0.8781 | 0.873 | 0.8782 | 0.8828 | 0.8915 | 0.8966 | 0.9069 | 0.9145 | -0.5 | 2.9295 | 3.4088 | 2.1902 | 1.3563 | 0.886 | 0.44699 | 0.26438 | 0.12213 | 0.069527 | 0.044522 | 0.030847 | 0.022477
| 8 | 0.96897 |  | 30 | 0.9351 | 0.9351 | 0.9054 | 0.8985 | 0.9065 | 0.9105 | 0.9166 | 0.9195 | 0.9279 | 0.9293 | -1 | 0.67874 | 0.82587 | 0.80538 | 0.6774 | 0.54119 | 0.34002 | 0.22309 | 0.11269 | 0.066356 | 0.043234 | 0.030203 | 0.022111
| 10 | 0.9395 |  | 35 | 0.9499 | 0.9499 | 0.9273 | 0.9238 | 0.9262 | 0.9288 | 0.9349 | 0.9406 | 0.9443 | 0.9479 | -1.5 | 0.30236 | 0.34589 | 0.37154 | 0.35659 | 0.32129 | 0.24107 | 0.1765 | 0.099436 | 0.061526 | 0.041128 | 0.029171 | 0.02154
|  |  |  | 40 | 0.9614 | 0.9614 | 0.9458 | 0.9415 | 0.9433 | 0.9444 | 0.9505 | 0.9532 | 0.9546 | 0.9551 | -2 | 0.17266 | 0.19016 | 0.20673 | 0.20962 | 0.20119 | 0.16958 | 0.13582 | 0.085361 | 0.055838 | 0.038483 | 0.027761 | 0.020775
|  |  |  | 45 | 0.9704 | 0.9704 | 0.9586 | 0.9555 | 0.9576 | 0.9597 | 0.963 | 0.9639 | 0.9664 | 0.9678 | -3 | 0.079387 | 0.08433 | 0.089707 | 0.093408 | 0.094346 | 0.090001 | 0.080802 | 0.060386 | 0.044074 | 0.032461 | 0.024436 | 0.018764
|  |  |  | 50 | 0.9769 | 0.9769 | 0.9672 | 0.9655 | 0.9678 | 0.9684 | 0.9723 | 0.9748 | 0.9766 | 0.9758 | -4 | 0.046395 | 0.048194 | 0.05025 | 0.052097 | 0.053225 | 0.053307 | 0.050778 | 0.042429 | 0.033768 | 0.026551 | 0.020873 | 0.016537
|  |  |  | 55 | 0.984 | 0.984 | 0.9768 | 0.9746 | 0.975 | 0.9774 | 0.9813 | 0.9839 | 0.9826 | 0.982 | -5 | 0.030559 | 0.031382 | 0.032309 | 0.033233 | 0.033974 | 0.034525 | 0.033988 | 0.030431 | 0.025825 | 0.021342 | 0.017495 | 0.014332
|  |  |  | 60 | 0.9878 | 0.9878 | 0.983 | 0.9805 | 0.9831 | 0.9838 | 0.9869 | 0.9872 | 0.9885 | 0.9874 | -6 | 0.02166 | 0.022105 | 0.022593 | 0.023054 | 0.023515 | 0.024002 | 0.023956 | 0.022396 | 0.019883 | 0.017131 | 0.014508 | 0.012267
|  |  |  | 65 | 0.9919 | 0.9919 | 0.9866 | 0.986 | 0.9882 | 0.9877 | 0.9921 | 0.9919 | 0.9937 | 0.9925 | -7 | 0.016157 | 0.016409 | 0.016709 | 0.016959 | 0.01721 | 0.017576 | 0.017625 | 0.016909 | 0.015501 | 0.013792 | 0.01207 | 0.010427
|  |  |  | 70 | 0.9949 | 0.9949 | 0.9922 | 0.9921 | 0.9928 | 0.9925 | 0.9974 | 0.9974 | 0.9973 | 0.9952
|  |  |  | 75 | 0.9966 | 0.9966 | 0.9941 | 0.9959 | 0.9955 | 0.9957 | 0.9994 | 0.998 | 0.9992 | 0.9983
|  |  |  | 80 | 0.9981 | 0.9981 | 0.9972 | 0.9987 | 0.9976 | 0.9973 | 1.0023 | 1.0027 | 1.0018 | 0.9989
|  |  |  | 85 | 0.9986 | 0.9986 | 1.0007 | 0.9966 | 0.9979 | 0.9981 | 1.0036 | 1.0033 | 1.0004 | 1.0005
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 95 | 0.9997 | 0.9997 | 0.9989 | 0.9975 | 0.9978 | 0.997 | 1.0022 | 1.0006 | 1.0008 | 0.9998
|  |  |  | 100 | 0.9982 | 0.9982 | 0.9976 | 0.9963 | 0.997 | 0.9972 | 1.0003 | 1.0016 | 0.9993 | 0.999
|  |  |  | 105 | 0.9968 | 0.9968 | 0.996 | 0.9948 | 0.9953 | 0.9955 | 1.0001 | 0.9995 | 1.0004 | 0.9955
|  |  |  | 110 | 0.9949 | 0.9949 | 0.9921 | 0.9919 | 0.9937 | 0.9942 | 0.9961 | 0.9969 | 0.9935 | 0.9913
|  |  |  | 115 | 0.9922 | 0.9922 | 0.9901 | 0.9867 | 0.9877 | 0.9878 | 0.9937 | 0.9931 | 0.9919 | 0.9917
|  |  |  | 120 | 0.988 | 0.988 | 0.9834 | 0.9827 | 0.9837 | 0.9833 | 0.9863 | 0.9872 | 0.9869 | 0.985
|  |  |  | 125 | 0.9839 | 0.9839 | 0.9767 | 0.9754 | 0.9764 | 0.9764 | 0.9815 | 0.9829 | 0.9825 | 0.9822
|  |  |  | 130 | 0.9779 | 0.9779 | 0.9686 | 0.964 | 0.9684 | 0.9695 | 0.9738 | 0.9756 | 0.9756 | 0.9755
|  |  |  | 135 | 0.9708 | 0.9708 | 0.958 | 0.9551 | 0.9571 | 0.9575 | 0.9629 | 0.9642 | 0.9651 | 0.9649
|  |  |  | 140 | 0.9621 | 0.9621 | 0.945 | 0.9401 | 0.9429 | 0.9444 | 0.9502 | 0.953 | 0.9556 | 0.9586
|  |  |  | 145 | 0.9512 | 0.9512 | 0.9284 | 0.9235 | 0.9267 | 0.9289 | 0.9357 | 0.9373 | 0.9418 | 0.9444
|  |  |  | 150 | 0.9357 | 0.9357 | 0.9064 | 0.8988 | 0.9029 | 0.9071 | 0.9148 | 0.9184 | 0.9248 | 0.9291
|  |  |  | 155 | 0.9136 | 0.9136 | 0.8756 | 0.868 | 0.8754 | 0.8813 | 0.8904 | 0.896 | 0.9025 | 0.9091
|  |  |  | 160 | 0.882 | 0.882 | 0.8361 | 0.8262 | 0.8355 | 0.8435 | 0.8548 | 0.8654 | 0.8776 | 0.887
|  |  |  | 165 | 0.7809 | 0.7809 | 0.7783 | 0.7731 | 0.7832 | 0.7952 | 0.8089 | 0.8202 | 0.8421 | 0.8568
|  |  |  | 168 | 0.73405 | 0.73405 | 0.7321 | 0.7282 | 0.7457 | 0.7612 | 0.7782 | 0.7914 | 0.8182 | 0.8371
|  |  |  | 170 | 0.6999 | 0.6999 | 0.6995 | 0.6987 | 0.7175 | 0.7354 | 0.7548 | 0.771 | 0.8005 | 0.8218
|  |  |  | 173 | 0.63448 | 0.63448 | 0.64045 | 0.6524 | 0.6763 | 0.698 | 0.72 | 0.7372 | 0.7734 | 0.798
|  |  |  | 175 | 0.60972 | 0.60972 | 0.61565 | 0.6275 | 0.6512 | 0.674 | 0.6982 | 0.7178 | 0.7574 | 0.7854
|  |  |  | 177 | 0.58602 | 0.58602 | 0.59205 | 0.6041 | 0.6282 | 0.6523 | 0.6785 | 0.7 | 0.744 | 0.774
|  |  |  | 178 | 0.5728 | 0.5728 | 0.5791 | 0.5917 | 0.6169 | 0.6421 | 0.6696 | 0.6916 | 0.7352 | 0.7656
|  |  |  | 179 | 0.56895 | 0.56895 | 0.5747 | 0.5862 | 0.6092 | 0.6322 | 0.6581 | 0.6817 | 0.7275 | 0.7611
|  |  |  | 180 | 0.5685 | 0.5685 | 0.5739 | 0.5847 | 0.6063 | 0.6279 | 0.6562 | 0.6789 | 0.7238 | 0.7574