## General Information
| General Information | Pd-103 - Theragenics model 200 (Theraseed)
Source Information
Source name: | Theragenics model 200 (Theraseed)
Radionuclide: | Pd-103
Model Number: | 200
Distributed by: | Theragenics Corporation
Manufactured by:
References
1) | "Update of AAPM Task Group No. 43 Report: A revised AAPM protocol for brachytherapy dose calculations"
Med Phys 31 (3), March 2004, 633-674
M.J. Rivard, B.M. Coursey, L.A. DeWerd, W.F. Hanson, M.S. Huq, G.S. Ibbott, M.G. Mitch, R. Nath,J.F. Williamson
Calculation details
Method: | Monte Carlo and TLD averaged
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised
2006-01-27

## Source design and materials
Source design and materials | Pd-103 - Theragenics model 200 (Theraseed)

## Geometry Function
Geometry Function |  |  |  | 103Pd - Theragenics model 200 (Theraseed)
Active source length,
| L = | 4.23 | mm

## Dose Rate Constant
Dose-Rate Constant |  |  | 103Pd - Theragenics model 200 (Theraseed)
L = | 0.686 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | 103Pd - Theragenics model 200 (Theraseed)
F(r,q)
r (cm)
q (º) | 0.25 | 0.5 | 0.75 | 1 | 2 | 3 | 4 | 5 | 7.5
0 | 0.619 | 0.694 | 0.601 | 0.541 | 0.526 | 0.504 | 0.497 | 0.513 | 0.547
1 | 0.617 | 0.689 | 0.597 | 0.549 | 0.492 | 0.505 | 0.513 | 0.533 | 0.58
2 | 0.618 | 0.674 | 0.574 | 0.534 | 0.514 | 0.517 | 0.524 | 0.538 | 0.568
3 | 0.62 | 0.642 | 0.577 | 0.538 | 0.506 | 0.509 | 0.519 | 0.532 | 0.57
5 | 0.617 | 0.6 | 0.54 | 0.51 | 0.499 | 0.508 | 0.514 | 0.531 | 0.571
7 | 0.579 | 0.533 | 0.519 | 0.498 | 0.498 | 0.509 | 0.521 | 0.532 | 0.568
10 | 0.284 | 0.496 | 0.495 | 0.487 | 0.504 | 0.519 | 0.53 | 0.544 | 0.59
12 | 0.191 | 0.466 | 0.486 | 0.487 | 0.512 | 0.529 | 0.544 | 0.555 | 0.614
15 | 0.289 | 0.446 | 0.482 | 0.49 | 0.523 | 0.54 | 0.556 | 0.567 | 0.614
20 | 0.496 | 0.442 | 0.486 | 0.501 | 0.547 | 0.568 | 0.585 | 0.605 | 0.642
25 | 0.655 | 0.497 | 0.524 | 0.537 | 0.582 | 0.603 | 0.621 | 0.64 | 0.684
30 | 0.775 | 0.586 | 0.585 | 0.593 | 0.633 | 0.654 | 0.667 | 0.683 | 0.719
40 | 0.917 | 0.734 | 0.726 | 0.727 | 0.75 | 0.766 | 0.778 | 0.784 | 0.82
50 | 0.945 | 0.837 | 0.831 | 0.834 | 0.853 | 0.869 | 0.881 | 0.886 | 0.912
60 | 0.976 | 0.906 | 0.907 | 0.912 | 0.931 | 0.942 | 0.96 | 0.964 | 0.974
70 | 0.981 | 0.929 | 0.954 | 0.964 | 0.989 | 1.001 | 1.008 | 1.004 | 1.011
75 | 0.947 | 0.938 | 0.961 | 0.978 | 1.006 | 1.021 | 1.029 | 1.024 | 1.033
80 | 0.992 | 0.955 | 0.959 | 0.972 | 1.017 | 1.035 | 1.046 | 1.037 | 1.043
85 | 1.007 | 0.973 | 0.96 | 0.982 | 0.998 | 1.03 | 1.041 | 1.036 | 1.043

## Radial Dose Function
Radial Dose Function Table |  |  | 103Pd - Theragenics model 200 (Theraseed)
Line source aproximation,
L = | 4.23 | mm | Point source aproximation,
r (cm) | g(r) |  | r (cm) | g(r)
0.1 | 0.911 |  | 0.1 | 0.494
0.15 | 1.21 |  | 0.15 | 0.831
0.25 | 1.37 |  | 0.25 | 1.154
0.3 | 1.38 |  | 0.3 | 1.22
0.4 | 1.36 |  | 0.4 | 1.269
0.5 | 1.3 |  | 0.5 | 1.248
0.75 | 1.15 |  | 0.75 | 1.137
1 | 1 |  | 1 | 1
1.5 | 0.749 |  | 1.5 | 0.755
2 | 0.555 |  | 2 | 0.561
2.5 | 0.41 |  | 2.5 | 0.415
3 | 0.302 |  | 3 | 0.306
3.5 | 0.223 |  | 3.5 | 0.226
4 | 0.163 |  | 4 | 0.165
5 | 0.0887 |  | 5 | 0.09
6 | 0.0482 |  | 6 | 0.0489
7 | 0.0262 |  | 7 | 0.0266
10 | 0.00615 |  | 10 | 0.00624

## Anisotropy Factor
Anisotropy Factor |  |  | 103Pd - Theragenics model 200 (Theraseed)
| r (cm) | fan(r)
| 0.25 | 1.13
| 0.5 | 0.88
| 0.75 | 0.859
| 1 | 0.855
| 2 | 0.87
| 3 | 0.884
| 4 | 0.895
| 5 | 0.897
| 7.5 | 0.918

## QA
| Dose rates for QA |  |  | Pd-103 - Theragenics model 200 (Theraseed)
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance using the 1-D formalism of Eq. (11) with interpolation for gL(r) and Øan(r). |  |  |  | QA along-away table obtained using TG-43U1S1 formalism and consensus dataset
| r (cm) | cGy h-1 U-1 | z / cm | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
| 0.5 | 3.014 | 0 | --- | 0.686 | 0.096228 | 0.02332 | 0.007085 | 0.0024683 | 0.00093163 | 0.00037209
| 1 | 0.587 | 0.5 | 3.0595 | 0.48501 | 0.088318 | 0.022941 | 0.0071448 | 0.0024971 | 0.00094821 | 0.00037884
| 1.5 | 0.199 | 1 | 0.39423 | 0.21832 | 0.063866 | 0.019251 | 0.00639 | 0.0023182 | 0.00089833 | 0.00036428
| 2 | 0.0837 | 1.5 | 0.12613 | 0.09058 | 0.040438 | 0.014622 | 0.0052714 | 0.0020218 | 0.0008095 | 0.00033519
| 3 | 0.0206 | 2 | 0.051378 | 0.04067 | 0.023882 | 0.010373 | 0.0041417 | 0.0016875 | 0.00069956 | 0.00029688
| 4 | 0.00634 | 2.5 | 0.023687 | 0.020237 | 0.013732 | 0.0069882 | 0.0031174 | 0.0013563 | 0.00058376 | 0.00025392
| 5 | 0.00221 | 3 | 0.011831 | 0.010779 | 0.0079001 | 0.0045051 | 0.0022491 | 0.0010541 | 0.00047689 | 0.0002118
| 6 | 0.000846 | 3.5 | 0.0063656 | 0.0060031 | 0.0045692 | 0.0029011 | 0.0015893 | 0.00079151 | 0.00037964 | 0.00017185
| 7 | 0.000342 | 4 | 0.0035344 | 0.0034393 | 0.0027344 | 0.00186 | 0.0010948 | 0.00058382 | 0.00029219 | 0.0001358
|  |  | 4.5 | 0.0020808 | 0.0020379 | 0.0016881 | 0.0012014 | 0.0007456 | 0.0004217 | 0.00021903 | 0.00010551
|  |  | 5 | 0.0012693 | 0.0012434 | 0.0010562 | 0.00077413 | 0.00050752 | 0.00030109 | 0.00016184 | 8.1396e-05
|  |  | 5.5 | 0.00078464 | 0.00077213 | 0.00066495 | 0.00050298 | 0.00034345 | 0.00021079 | 0.00011612 | 6.3262e-05
|  |  | 6 | 0.00049141 | 0.00048553 | 0.00042423 | 0.00033074 | 0.0002316 | 0.00014528 | 8.3228e-05 | 5.0636e-05
|  |  | 6.5 | 0.00031232 | 0.00030942 | 0.00027394 | 0.00021858 | 0.00015394 | 9.8547e-05 | 6.0779e-05 | 4.2719e-05
|  |  | 7 | 0.00020125 | 0.00019933 | 0.00017786 | 0.00014224 | 0.00010086 | 6.7918e-05 | 4.6541e-05 | 3.8943e-05