## General Information
| General Information | Pd-103 - NASI model MED3633
Source Information
Source name: | NASI model MED3633
Radionuclide: | Pd-103
Model Number: | MED3633
Distributed by: | North American Scientific Inc.
Manufactured by:
References
1) | "Update of AAPM Task Group No. 43 Report: A revised AAPM protocol for brachytherapy dose calculations"
Med Phys 31 (3), March 2004, 633-674
M.J. Rivard, B.M. Coursey, L.A. DeWerd, W.F. Hanson, M.S. Huq, G.S. Ibbott, M.G. Mitch, R. Nath,J.F. Williamson
Calculation details
Method: | Monte Carlo and TLD averaged
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised
2006-01-27

## Source design and materials
Source design and materials | Pd-103 - NASI model MED3633

## Geometry Function
Geometry Function |  |  |  | 103Pd -NASI model MED3633
Active source length,
| L = | 4.2 | mm

## Dose Rate Constant
Dose-Rate Constant |  |  | 103Pd -NASI model MED3633
L = | 0.688 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | 103Pd -NASI model MED3633
F(r,q)
r (cm)
q (º) | 0.25 | 0.5 | 1 | 2 | 5 | 10
0 | 1.024 | 0.667 | 0.566 | 0.589 | 0.609 | 0.733
10 | 0.888 | 0.581 | 0.536 | 0.536 | 0.569 | 0.641
20 | 0.85 | 0.627 | 0.603 | 0.614 | 0.652 | 0.716
30 | 0.892 | 0.748 | 0.729 | 0.734 | 0.756 | 0.786
40 | 0.931 | 0.838 | 0.821 | 0.824 | 0.837 | 0.853
50 | 0.952 | 0.897 | 0.89 | 0.891 | 0.901 | 0.905
60 | 0.971 | 0.942 | 0.942 | 0.94 | 0.948 | 0.939
70 | 0.995 | 0.976 | 0.974 | 0.973 | 0.98 | 0.974
80 | 1.003 | 0.994 | 0.997 | 0.994 | 1 | 0.986

## Radial Dose Function
Radial-Dose Function Table |  |  | 103Pd -NASI model MED3633
Line source aproximation,
L = | 4.2 | mm |  | Point source aproximation,
r (cm) | g(r) |  |  | r (cm) | g(r)
0.25 | 1.331 |  |  | 0.25 | 1.123
0.3 | 1.322 |  |  | 0.3 | 1.17
0.4 | 1.286 |  |  | 0.4 | 1.201
0.5 | 1.243 |  |  | 0.5 | 1.194
0.75 | 1.125 |  |  | 0.75 | 1.113
1 | 1 |  |  | 1 | 1
1.5 | 0.77 |  |  | 1.5 | 0.776
2 | 0.583 |  |  | 2 | 0.589
2.5 | 0.438 |  |  | 2.5 | 0.443
3 | 0.325 |  |  | 3 | 0.329
3.5 | 0.241 |  |  | 3.5 | 0.244
4 | 0.177 |  |  | 4 | 0.179
5 | 0.098 |  |  | 5 | 0.099
6 | 0.053 |  |  | 6 | 0.054
7 | 0.028 |  |  | 7 | 0.028

## Anisotropy Factor
Anisotropy Factor |  |  | 103Pd -NASI model MED3633
| r (cm) | fan(r)
| 0.25 | 1.257
| 0.5 | 0.962
| 1 | 0.903
| 2 | 0.895
| 5 | 0.898
| 10 | 0.917

## QA
Dose rates for QA |  |  | Pd-103 - NASI model MED3633
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources using
the 1-D formalism of Eq. (11) with interpolation for gL(r) and Øan(r).
| r (cm) | fan(r)
| 0.5 | 3.184
| 1 | 0.626
| 1.5 | 0.215
| 2 | 0.0914
| 3 | 0.0227
| 4 | 0.00697
| 5 | 0.00247
| 6 | 0.000933
| 7 | 0.000364